EPHYS INC - AUTOMATED ANALYSIS NOTES FOR YIANNOS

2-2-26_DMG

SOFTWARE

Clampex 11.4.3 (for data acquisition)

Clampfit 11.4.3 (for analysis))

Axon™ pCLAMP™ 11 Electrophysiology Data Acquisition & Analysis Software Download Page

DATA ORGANIZATION

Project Folder\Animal ID Folder\Sample # Folder\list of ABF files

There are 11 total protocols executed during a single recording, with each protocol producing a different ABF file. In a Sample # folder there may be more than 11 ABF files due to multiple attempts, or there may be fewer than 11 files that are still in need of analysis. Each ABF file name contains the collection date, the initials of the recorder, and the file order number starting with 0000 (e.g., 0000-0009 for 10 total files). Here is an example of an ABF file name: 2026_01_16_DMG_0037. Although each protocol is recorded in the same order, due to inconsistencies each ABF file requires opening in Clampfit to visually inspect the associated traces and confirm it contains the appropriate protocol data.

The two folders provided are two Animal ID Folders (64280 & 64281), each containing several Sample # Folders that have ABF files to work and practice with. Each of the Sample # Folders have exactly 11 ABF files, all in the protocol order detailed below, except for Sample # Folder 1670 (this folder has 12 files).

PROTOCOL OVERVIEW

Each protocol executed during a recording is associated with two types of signals that Clampfit recognizes and plots, the response signal obtained from the cell (mV) and the stimulus signal used to manipulate the cell (pA). The stimulus signals for a given protocol will vary as they depend upon the different protocol parameters. Stimuli may be presented in very brief pulses (e.g., 2 ms) or longer steps ( e.g., 1000 ms) and are usually presented in escalating fashion (e.g., increasing 25 pA each successive step) or from a list (e.g., 0, 100, 500, 700, 1000, 500, 100 pA, etc.). Stimuli trains within a protocol are also usually repeated in sweeps, which are also enumerated below when appropriate.

LIST OF PROTOCOLS & PARAMETERS

- EPSC – only voltage clamp measure ( 3 sec sEPSC @ -67 mV )

- RMP – 15 s with no holding current applied

- Membrane Test – 1000 sweeps (25 ms step, -50 pA)

- Ri/Ih – 20 sweeps (1 s step, starting at 0.3 s) from list (-200, -100, -200, -100, -50, -40, -30, -20, -10,

0, -50, -40, -30, -20, -10, 0, 25, 50, 25, 50)

- AP coarse – 12 sweeps (2 ms pulse) from list (0,150,300,400,500,600,700,800,850,900,950,1000 pA)

- AP fine – 10 sweeps (2 ms pulse, 25 pA) starting at 300 pA (user defined)

- Ramp – 6 second ramp (linear increase of 620 pA from holding current)

- 5 AHP – 2 sweeps (5 pulses each: 1200 pA, 2 ms, 100Hz), held at -60mV

- 25 AHP – 2 sweeps (25 pulses each: 1200 pA, 2 ms, 100Hz), held at -60mV

- 1s AP – 15 sweeps (1000 ms steps, starting at 3000 ms); first 10 sweeps: 50 pA each, from 0 to 500 pA; 5 final sweeps: 100 pA each, from 500 to 1000 pA)

- 50 AHP – 2 sweeps (50 pulses each: 2000 pA, 2 ms, 50Hz), held at -67mV

PROTOCOL DATA DETAILS

- EPSC – DOESN’T NEED ATTENTION AT THIS TIME

- RMP:

Clampfit – we obtain 1 metric, the mean value within a 200 ms window starting at t=0 ms for final RMP value

- Membrane Test – NOT APPLICABLE

- Ri/Ih:

3 Final Metrics = Input Resistance (MOhm), Sag Amp (mV), Sag Ratio

Clampfit – we obtain 3 measurements from each sweep (BASELINE: mean value within 200 ms window, starting at t=0 ms; PEAK: peak negative amplitude within 200 ms window, starting 300 ms (when stimulus begins); STEADY STATE: mean value within 200 ms window, starting at 1000 ms), each containing 2 different types of data - the response signal from the cell (mV) and the stimulus used for the given sweep (pA), which fluctuates according to the list (above).

Excel – For the response signal (mV), we subtract BASELINE from both PEAK and STEADY STATE. For the stimulus signal (pA), we subtract BASELINE from STEADY STATE. We then calculate final Input Resistance value using the slope of the line resulting from plotting STEADY STATE (mV) vs STEADY STATE (pA), only plotting sweeps between -50 and +50 pA. We also calculate final Sag Amplitude value by subtracting the BASELINE response signal (mV) from the PEAK stimulus signal (pA). And finally, we calculate Sag Ratio as a response signal (mV). Sag Ratio = (STEADY STATE – BASELINE) / (PEAK – BASELINE).

- AP coarse: The protocol data from this file does not need to be use unless there is no AP fine protocol file.

- AP fine:

7 Final Metrics = Peak Amp (mV); Half Width (ms); Max Rise Slope (mV/ms); Max Decay Slope (mV/ms); sweep #; stimulus signal value associated with sweep # (Current Threshold, pA); Voltage Threshold (mV)

Quality Control Filter = RMP<VThresh. We want to apply this filter at the beginning of the script, if possible, to sort recordings based on quality and restrict further analysis to good recordings. Ideally, we’d like the user to visualize two different plots (within the same window, detailed below) for each event/spike to confirm final acceptance of the recording to continue with the analysis.

Voltage Threshold (mV) & RMP<VThresh (boolean): The Voltage Threshold is calculated by transferring the ABF response signal (mV) data within a defined window (t=0.2975 to 0.3015 ms) to create a phase plot (dV/dt vs mV) and find the point on the x-axis that corresponds to the first value > 28. The RMP<VThresh is a Boolean filter (if Voltage Threshold > RMP, accept recording and continue with analysis; if Voltage Threshold < RMP, reject recording and place in a RMP>VThresh folder for user review. After Voltage Threshold is calculated and before RMP<VThresh is, we’d like the user to visualize both the phase plot (dV/dt vs mV) and the response signal (mV) for every event/spike, with each plot displaying the calculated Voltage Threshold represented as a v-line (dotted line on x-axis). This will enable the user to either accept or reject the calculated Voltage Threshold value for further analysis (Boolean command) and sort rejected recordings into a User Rejected VThresh folder.

Clampfit – we obtain 5 measurements (BASELINE: mean value within 200 ms window starting      at t=0 ms; AND THE FOLLOWING 4 MEASURES ARE ALL TAKEN FROM THE SAME WINDOW (t=0.29 s to 0.31 s): Peak Amp (mV); Half Width (ms); Max Rise Slope (mV/ms); Max Decay Slope (mV/ms). Each of these is taken only from the sweep during which the first event/spike occurred, and the sweep # as well as the stimulus signal value/Current Threshold (pA) associated with the sweep # should be collected along with other metrics.

Excel – Final Peak Amp is calculated by subtracting BASELINE from Peak Amp above. Final values for other metrics are same as those obtained from Clampfit.

- Ramp:

Final Metrics = Rheobase (pA) & Rheobase Step Up Firing Frequency (pA)

Clampfit – We obtain both the peak positive and negative peak amplitude of both the stimulus signal (pA) and the response signal (mV) within a user-defined window: from t=0 ms to the very beginning of first response signal spike (mV).

Excel – Calculate the Rheobase by subtracting the negative peak from positive peak of the stimulus response (pA).

- 5 AHP:

Final Metrics = Peak, 200ms & 1000ms (mV)

Clampfit – we obtain 4 measurements from each sweep (BASELINE: mean value within 200 ms window, starting @ t=0 ms; PEAK: peak negative amplitude within 200 ms window, starting 2 ms after last of 5 stimulus pulses; 200ms: mean value within 2 ms window, starting 200ms after final stimulus pulse; 1000ms: mean value within 200 ms window, starting 1000 ms after final stimulus pulse).

Excel – subtract BASELINE from PEAK, 200ms, & 1000ms values, and finally average the 2 values of each measurement obtained from both sweeps for final PEAK, 200ms & 1000ms values

- 25 AHP:

Final Metrics = Peak, 200ms & 1000ms (mV)

Clampfit – we obtain 4 measurements from each sweep (BASELINE: mean value within 200 ms window, starting @ t=0 ms; PEAK: peak magnitude within 200 ms window, starting 2 ms after last of 25 stimulus pulses; 200ms: mean value within 2 ms window, starting 200ms after final stimulus pulse; 1000ms: mean value within 200 ms window, starting 1000 ms after final stimulus pulse).

Excel – subtract BASELINE from PEAK, 200ms, & 1000ms values, and finally average the 2 values of each measurement obtained from both sweeps for final PEAK, 200ms & 1000ms values

- 1s AP:

Final Metrics = firing frequency (Hz), inter-spike interval (ms), First FF (Hz), & Rheobase SU FF (Hz)

Clampfit – Sets a threshold of 0 mV to filter events (only spikes that exceed this threshold from more negative baselines) and we want to add a second filtration criteria: only include spikes within sweeps that contain 50 spikes or less (reject sweeps that contain more than 50 spikes and set into reject folder for user review). We then take the total number of events/spikes per sweep and the total net events per recording.

Excel – Calculate firing frequency (Hz) and inter-spike interval (ms) per each sweep and total recording. Identify firing frequency during the sweep for which the first spikes occurred (First FF). Identify the firing frequency during the first sweep with a corresponding stimulus signal larger than the Rheobase value calculated earlier (Rheobase SU FF).

- 50 AHP:

Final Metrics = Peak, 200ms & 1000ms (mV)

Clampfit – we obtain 4 measurements from each sweep (BASELINE: mean value within 200 ms window, starting @ t=0 ms; PEAK: peak magnitude within 200 ms window, starting 2 ms after last of 50 stimulus pulses; 200ms: mean value within 2 ms window, starting 200ms after final stimulus pulse; 1000ms: mean value within 200 ms window, starting 1000 ms after final stimulus pulse).

Excel – subtract BASELINE from PEAK, 200ms, & 1000ms values, and finally average the 2 values of each measurement obtained from both sweeps for final PEAK, 200ms & 1000ms values

SCRIPT IDEAS

We’d like to start the analysis with a quality control filtration process, calculating the final metrics RMP and Voltage Threshold first to enable the RMP<VThresh Boolean filter. This should sort all the disqualified recordings into a RMP>VThresh folder for further review and allow the recordings that reach criteria to continue for further analysis. Thus, the order in which the protocols should be analyzed is as follows:

RMP → AP fine (filters) → Ri/Ih → Ramp → 1s AP → 5 AHP → 25 AHP → 50 AHP

If possible, we’d like to add nominal variables before the data is exported to further sort for user convenience (e.g., SEX: male/female; BEHAVIOR: naïve/CFM). This may be done by assigning each Sample # their appropriate variable, as the Sample # is part of original file organization. Also, we’d like to name and add, or subtract nominal variables at will. The final data should be exported to 1 of 3 folders based on quality control results (PASSED, User Rejected VThresh, RMP>VThresh), each containing an Excel file that contains a data table with the selected nominal variables and the appropriate metrics for the given folder produced by the script. The file in the PASSED folder should contain data for every final metric, with each protocol assigned its own individual tab within the same Excel file, and each tab containing that protocol’s final metrics. Each tab’s data table should also repeatedly display the Sample # and nominal variables corresponding to each listed final metric (all data should be listed in order of Sample #). The User Rejected VThresh data table should contain the file path, Sample # and sweep # associated with each rejected calculation, along with the corresponding nominal data. Similarly, the RMP>VThresh data table should contain the file path, Sample # and nominal data of each rejected recording.